Project#1 Spectrum Analyzer Interface Standard:
Command Structure Rev 2: 28 September 2021

 To provide a unified set of commands for Project 1 Spectrum Analyzer utilizing USB, along with support information to allow device programming of the Project.

Description: The Project 1 Spectrum Analyzer RF Hardware is primarily locally controlled using Arduino-based controllers (but not limited to). A PC or Laptop is typically the HID for the use of the Spectrum Analyzer and the interface hardware is a USB connection. The interface specification described here does not limit the controllers to Arduino nor does it limit the connection to USB. The use of Ethernet or WiFi could be implemented with it. So in this Interface Standard, the term Arduino could be replaced with Raspberry Pi, etc and the USB Bus could be replaced with Ethernet, etc. Also, this Interface Standard is compatible with both RF Assemblies 2.0.0.00 and 3.0.0.00

Background: Various Spectrum Analyzer RF components require serial control. The majority of these are PLL Synthesizers with 32-bit registers. One exception is a digital attenuator with a 7-bit serial register, and a provisional 12-bit Analog to Digital converter. Within the 32-bit registers there are multiple counters that require programming. All of these components work with a 3-wire (or 4) SPI bus. Three lines are normally used: Clock, Data Line (Arduino >> RF Hardware) and Latch Enable. A fourth line has been provisioned for the use of ADC hardware, which is a Data Line (RF Hardware >> Arduino). Therefore all RF Hardware is accommodated by multiple 32-bit words. In the same way, all PC/Laptop communications is handled exactly the same way, but not necessarily with the exact same data in the words. For the dual purposes of addressing RF Hardware’s multiple devices and greatly improved sweep rates, a command structure allows for a single 32-bit word on the USB Bus to issue multiple 32-bit words from Arduino to RF Hardware.

To make all this practical for Arduino, it will be used only for communication, control and translation. For purposes of throughput, no Floating Point calculations will be needed, as all math can be done with Boolean and Integer operations to enhance speed. All math operation requiring Floating Point operations are to be done in the PC/Laptop. The Arduino can efficiently perform bit masking and bit-shifting.

Method: A command word is issued with a unique code not common to any of the 32-bit words used to control the RF Hardware. None of the RF Hardware sees 0x yy yy yy FF, as a valid command. This allows for multiple commands directed to different addresses that can be issued from the PC/Laptop to the Arduino that will be uniquely seen as a command as well as embedded data further enhance throughput and hence, sweep rate, as will be seen. Any 32-bit word that does not have ‘FF’ as its LSB’s will be seen as data, either abbreviated into shorter form or direct (unmodified), the function of which was established by the most recent command word. When a new command word is issued by the PC/Laptop, the Arduino translation is changed accordingly to suit the command. It can do so by calling different subroutines to parse the 32-bit word.

This document details programming of various Frequency Counters used in the LO’s. It does not detail the other functions, such as power control, phase, and many other register capabilities. Not all these command words need to be implemented initially. Certain proposed commands are listed as provisional,
and not all of these commands need to be implemented at any time.

Outline

1. Command Structure.

2. Specific Commands.

2.1 Digital Attenuator U103

2.2 LO1 ADF4356 U100

2.3 LO2 (U203) and LO3 (U307) MAX2871

2.3 a) Single Word Embedded Data and Specific Command for Frequency Control.

2.3 b) Block Transfers.

2.3 c) Data Words.

2.4 Reference Clock Control.

2.5 Analog to Digital (ADC) Converter Data Return

2.6 PLL_MUX Specific Command

3. General LO Commands

APPENDIX A. Fractional Synthesis Calculation with MAX2871

APPENDIX B. Sample Program in FreeBasic.

APPENDIX C. Table of Default Settings

APPENDIX D. Default Arduino Code
1. Command Structure:
Embedded Data (as required)					Command					Address			Command Flag 0xFF
31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
D	D	D	D	D	D	D	D	D	D	D	D	D	D	D	D	C	C	C	C	C	A	A	A	1	1	1	1	1	1	1	1

Note: 4-Bit Nibble Groups are shown in these diagrams for easier Hexadecimal Coding.

D= Data (Embedded) 16 bits max <31:16>.

C=Command Code. 5 Bits <15:12>

A= Address Code 3 Bits <10:8>
      Ex: 0h=Digital Attenuator; 1h=LO1; 2h=LO2; 3h=LO3; 4h=Reference Frequency Control;
        5h=PLL_MUX poll.

For ALL Command Words (Specific and General) Bits <7:0> = FFh Command Flag

There are two types of Device Commands, General and Specific. General Commands support total register control, whereas Specific provide a much faster control method, but relies on previously Arduino programmed default values (Refer to Appendix C).

2. Specific Commands. This Command Method is preferred for most programming.

2.1 Digital Attenuator U103: Address 0 Embedded Data and Command. A type of Specific Command.
  Address=0h, Command=0h

Embedded Data (as required)	Command	Address	Command Flag 0xFF

31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
-	-	-	-	-	-	-	-	-	D6	D5	D4	D3	D2	D1	D0	0	0	0	0	0	0	0	0	1	1	1	1	1	1	1	1

Note: 4-Bit Nibble Groups are shown in these diagrams for easier Hexadecimal Coding.

 This Digital Attenuator Command Word has the Data embedded for attenuator control.
From this 32-bit word, bits <22:16> map by Arduino into the Digital Attenuator Data bits <D6:D0> as per the pSemi PE43711 Datasheet. The PE43711 is a 0-31.75dB Digital Attenuator with 0.25dB steps.
Ex: 0x 00 00 00 FF is 0dB; 0x 00 01 00 FF is 0.25dB…..0x 00 7F 00 FF is 31.25 dB, etc

2.2 LO1 ADF4356 U100 Single Word Embedded Data and Command for Large Steps (faster).
For implementing Integer-N operation; which allows large steps at Fpfd spacing (typically ≥ 30 MHz),
using only one 32-bit word. The value of Fpfd depends on Fref, the Reference Clock Frequency (60 or 100 MHz) and the overall R divide value, typically set at 2. This method provides a faster method of programming LO1 after the initialization procedure at power-up by Arduino or by use of General LO Commands in section 3.

NOTE: For LO1 Programming, there is no data word returned automatically from the ADC’s. That is done after LO2 or LO3 is programmed. See Section 2.3.

Integer Mode:
LO1 [MHz] = (Fref[MHz] /R)* (N) (Valid for 3400 ≤ LO1[MHz] ≤ 6800)
then: N= LO1[MHz] *R / Fref[MHz]

So for LO1= 3600 MHz (the low end of our LO1 band) with Fref=60MHz, R=2, then N=120

For LO1=3630 then N=121 and so on... to LO1=6630MHz N=221 MHz.
If the Reference Frequency is changed to 100 MHz, then with R=2, N needs to be recalculated.

The Full use of the ADF4356 16-bit wide Integer counter is allowed in this Specific Command, although for our purposes values of N=72d (48h) to N=221d (DDh) are all that will be needed.

LO1: Address =1h Command=(see Spreadsheet List)

Embedded Data (as required)	Command	Address	Command Flag 0xFF

31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
N / 16	N / 15	N / 14	N / 13	N / 12	N / 11	N / 10	N / 9	N / 8	N / 7	N / 6	N / 5	N / 4	N / 3	N / 2	N / 1	C	C	C	C	C	0	0	1	1	1	1	1	1	1	1	1

The values in <31:16> are directly mapped by Arduino in LO1’s (ADF4356) Register 0 at <N16:N1>.
Note: For reasons I cannot comprehend, ADI uses sub-register bits as N16:N1 which is starts at N1, but all the register bits are DB31:DB0 which start at DB0. Fortunately, Maxim is consistent.

Example 1: Fref=60 MHz R=2(default) we want LO1=3600MHz, from above equation N=120d =78h
  Then the Specific Command for LO1 is 0x 00 78 11 FF

 Example 2: Fref=60 MHz R=2(default) we want LO1=5400MHz, from above equation N=180d =B4h
  Then the Specific Command for LO1 is 0x 00 B4 11 FF

Example 3: Fref=100 MHz R=2(default) we want LO1=4700MHz, from above equation N=94d =5Eh
  Then the Specific Command for LO1 is 0x 00 5E 11 FF.

NOTE: Refer to Section 2.4 for proper use of the reference clock.

2.3 LO2 (U203) and LO3 (U307) MAX2871
2.3 a) Single Word Embedded Data and Specific Command for Frequency Control.
      Refer to MAX2871 Datasheet for all register details.
Medium and Fine frequency control with full control of integer, fractional and modulus counters can be done by programming LO2 (and LO3 for 45 MHz) which can also provide much better sweep rates.
This method provides a faster method of programming LO2 and LO3 after the initialization procedure at power-up by Arduino or by use of General LO Commands in section 3.
NOTE: A single Specific Command word is issued to instruct the Arduino as to which LO (LO2 or LO3) is to be programmed.

LO2: Address =2h For Command see Spreadsheet List
Embedded Data (as required)	Command	Address	Command Flag 0xFF

31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
-	-	-	-	-	-	-	-	-	-	-	-	BT3	BT2	BT1	BT0	C	C	C	C	C	0	1	0	1	1	1	1	1	1	1	1

LO3: Address =3h For Command see Spreadsheet List
Embedded Data (as required)	Command	Address	Command Flag 0xFF

31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
-	-	-	-	-	-	-	-	-	-	-	-	BT3	BT2	BT1	BT0	C	C	C	C	C	0	1	1	1	1	1	1	1	1	1	1

2.3 b) Block Transfers. This word’s bits <19:16> or <BT3:BT0> denote the number of Data Words to be sent to the Arduino per Block Transfers. For example 1h denotes 1 32-bit Data Word sent to the Arduino which then invokes the 16-bit return of ADC data (refer to section 2.5). Likewise, 8h will denote (8) successive 32-bit contiguous Data Words sent to the Arduino, followed by (8) 16-bit ADC Data Words returned. This has been observed as a very efficient method of data transfer at 1 Megabaud rates. Only (1) Specific Command Word is needed to invoke as many block transfers as desired, as long as each block has the same number of 32-bit Data Words/ 16-bit ADC words as any other block.

2.3 c) Data Words. Data Words are sent to LO2 or LO3. There is no limit on the number of these Data Words, as long as they are preceded by the appropriate command from a). This can allow for rapid sweeping of LO2 (and LO3 for 45 MHz fine steps), while allow LO1 to stay mostly stationary. Direct Word <31:0> maps directly into any of the LO’s <DB31:DB0> registers, as directed by the command word. The value of <7:0> must never equal 0xFF, as this signifies a command word, and the LO’s don’t see this as a valid value.

FRACTIONAL COUNTER(F) 12 Bits					MODULUS (M) COUNTER 12 BITS					N COUNTER 8 BITS
31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
F / 11	F / 10	F / 9	F / 8	F / 7	F / 6	F / 5	F / 4	F / 3	F / 2	F / 1	F / 0	M / 11	M / 10	M / 9	M / 8	M / 7	M / 6	M / 5	M / 4	M / 3	M / 2	M / 1	M / 0	N / 7	N / 6	N / 5	N / 4	N / 3	N / 2	N / 1	N / 0

The LO2 output frequency now is defined by the equation in Fractional Mode:
LO2[MHz]= (Fref[MHz]/R)*(N+F/M). The MAX2871 output divider is set to 1, so it does not affect
Refer to algorithm in Appendix A for N, F and M values.

LO2 Example 1: LO2 [MHz] to be 3915 MHz. Calculate N, F and M.
Results are N=130d =82h F=1d=1h M=2d=2h
 LO2[MHz]=(60MHz/2)(130+1/2)=3915 MHz

LO2 Example 2: LO2 [MHz] to be 4099.55 MHz. Calculate N, F and M.
Results are N=136d= 88h F=391d=187h M=600d=258h
 LO2[MHz]=(60MHz/2)(136+391/600)=4099.55 MHz
LO2 Example 3: LO2 [MHz] to be 4112.345 MHz. Calculate N, F and M.
Results are N=137d=89h F=295d=127h M=3774d=EBEh
What?? LO2[MHz]=(60MHz/2)(137+295/3774)=4112.344992 MHz!! An Error of 8 Hz!
Relax, it’s fractional synthesis, not perfectly exact, but close enough!. If less error is required, which could be necessary for the 45 MHz channel, then do the very fine steps with LO3, with LO2 held fixed.

 Now what to do with N, F and M? They can be easily mapped into the Data Words as the diagram shows. Convert values to Hexadecimal and map them as shown for LO2 Example 3.

        F=127h M=EBEh N=89h
FRACTIONAL COUNTER(F) 12 Bits					MODULUS (M) COUNTER 12 BITS					N COUNTER 8 BITS
31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
0	0	0	1	0	0	1	0	0	1	1	1	1	1	1	0	1	0	1	1	1	1	1	0	1	0	0	0	1	0	0	1

c) The LO3 output frequency also is defined by this equation in Fractional Mode:
LO3[MHz]= ((Fref[MHz]/R)*(N+F/M))/16 (the value 16 is due to the MAX2871 output divider)
Refer to algorithm in Appendix A for N, F and M values.

LO3 Example 1: LO3 [MHz] to be 270 MHz
Results are N=144d=90h F=0d=0h M=2d=2h
LO3[MHz]= (60MHz/2)(144+0/2)/16 =270.00000MHz

LO3 Example 2: LO3 [MHz] to be 360 MHz
Results are N=192d=C0h F=0d=0h M=2d=2h
LO3[MHz]= (60MHz/2)(192+0/2)/16 =360.00000MHz

LO3 Example 3: LO3 [MHz] to be 270.001 MHz
Results are N=144d=90h F=1d=1h M=1875d=753h
LO3[MHz]= (60MHz/2)(144+1/1875)/16 =270.00100MHz

 Now what to do with N, F and M? They can be easily mapped into the Data Words as the diagram shows. Convert values to Hexadecimal and map them as shown for LO3 Example 3.

        F=1h M=753h N=90h
FRACTIONAL COUNTER(F) 12 Bits					MODULUS (M) COUNTER 12 BITS					N COUNTER 8 BITS
31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
0	0	0	0	0	0	0	0	0	0	0	1	0	1	1	1	0	1	0	1	0	0	1	1	1	0	0	1	0	0	0	0

2.4 Reference Clock Control. Address=4h, Command=(see Spreadsheet List)

Embedded Data (as required)	Command	Address	Command Flag 0xFF

31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
-	-	-	-	-	-	-	-	-	-	-	-	-	-	-	-	C	C	C	C	C	1	0	0	1	1	1	1	1	1	1	1

Note: 4-Bit Nibble Groups are shown in these diagrams for easier Hexadecimal Coding.

The Spreadsheet has the listing for how to command REF 1, REF 2 or No Reference. The selected reference is used by all 3 LO’s (2 LO’s if 45 MHz is disabled).
Default Reference is 60 MHz (U500). Alternate Reference of 100 MHz (U501) is used primarily when the RF frequency being received is 60 MHz or a harmonic of 60 MHz, or if the RF frequency results in a signal that falls within the spectrum analyzer’s IF bandwidth.
A “B.I.T.E.” Test can be achieved by disabling both references and look for “UNLOCK” on all three LO’s. This could be useful at power-up. See PLL_MUX command.

To avoid 30 MHz Phase-Frequency Detector (PFD) Harmonics that are present with the 60 MHz reference clock, use the 100MHz reference clock which has 50MHz (PFD) harmonics. Likewise, the reverse is true. Try to use the 60 MHz reference clock since its phase jitter is >3 times better than the 100 MHz reference clock.

2.5 Analog to Digital (ADC) Converter Data Return. (No Command Word required)
The ADC value from the log amps do not require a separate command word to invoke. Rather, the type of an LO2 or LO3 Specific Command invokes the return of a 16-bit integer(s) automatically.
To command the Arduino to automatically send back either a 315 MHz ADC value or a 45 MHz ADC
value after either LO2or LO3 is commanded by a Specific Command.

a) If LO3 is enabled at any power level, the 45 MHz Log amp ADC Value is returned.
b) If LO3 is disabled, then the 315 MHz Log Amp ADC Value is returned.
The Arduino always knows the status of LO3. For example if Command 1h was sent to LO3’s address (3h) then Arduino knows LO3 is disabled. Otherwise LO3 is enabled. Also, reading 315 MHz ADC data when LO3 is enabled is not optimal due to LO3 signal blocking.

The resolution could be an integer word (16 bit), which could originate from either the 10-bit Arduino Analog inputs or the 12-bit Digitizer board. The Arduino program is to be written for either -type of input (i.e. an Arduino program could handle its own 10-bit Analog inputs or be written to handle the 12-bit Digitizer which needs additional commands to function). A more elaborate Arduino program could handle either input by querying the AD7887, if desired via the SPI_MISO line and /SEL_ADC, then knowing that the Digitizer board is functional, the Arduino program can use the AD7887 values in place of its own Analog Inputs.

2.6 PLL_MUX Specific Command. Address=5h
Even though it is not a physical device, the PLL_MUX Specific Command will invoke the return of a 16-bit word which has the PLL Lock status. It requires the Arduino to program the respective register that enables that PLL_MUX line of that LO, one at a time, sense the state whether high(lock) or low(unlock) and then disable that LO’s PLL_MUX line. A simple 16 bit (unsigned integer) word is returned after the Arduino polls all three LO’s. A logic ‘0’ indicates UNLOCK, a logic ‘1’ indicates LOCK for that specific LO:

15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
P	P	P	P	P	P	P	P	P	P	P	P	LO3	LO2	LO1	P

Note: P=Provisional; 4-Bit Nibble Groups are shown in these diagrams for easier Hexadecimal Coding.

3. General LO Commands (Provisional) for LO1, LO2 or LO3. This permits total and direct control of any LO, any register at the price of speed. Specific Commands in section 2 are faster at cost of flexibility. These Command Words are used to direct the Arduino on what to do with the following Direct Data Words. After a Command word is issued (with the Command Flag 0xFF) any number of Direct Data Words can follow. They are to be directed to the specific LO until a different Command Flag is seen.

LO1 Command Word. Note: Address <11:8> =1h, Command=0h

Embedded Data (as required)	Command	Address	Command Flag 0xFF

31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
-	-	-	-	-	-	-	-	-	-	-	-	-	-	-	-	0	0	0	0	0	0	0	1	1	1	1	1	1	1	1	1

LO2 Command Word. Note: Address <11:8> =2h, Command=0h

Embedded Data (as required)	Command	Address	Command Flag 0xFF

31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
-	-	-	-	-	-	-	-	-	-	-	-	-	-	-	-	0	0	0	0	0	0	1	0	1	1	1	1	1	1	1	1

LO3 Command Word. Note: Address <11:8> =3h, Command=0h

Embedded Data (as required)	Command	Address	Command Flag 0xFF

31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
-	-	-	-	-	-	-	-	-	-	-	-	-	-	-	-	0	0	0	0	0	0	1	1	1	1	1	1	1	1	1	1

Follow the Command Word for the LO being programmed with (N) Direct Data Words:
There is no limit on number of direct data words. This sequence is terminated at the next command word.
LO1/LO2/LO3 DIRECT DATA BITS
31	30	29	28	27	26	25	24	23	22	21	20	19	18	17	16	15	14	13	12	11	10	9	8	7	6	5	4	3	2	1	0
DB / 31	DB / 30	DB / 29	DB / 28	DB / 27	DB / 26	DB / 25	DB / 24	DB / 23	DB / 22	DB / 21	DB / 20	DB / 19	DB / 18	DB / 17	DB / 16	DB / 15	DB / 14	DB / 13	DB / 12	DB / 11	DB / 10	DB / 9	DB / 8	DB / 7	DB / 6	DB / 5	DB / 4	DB / 3	DB / 2	DB / 1	DB / 0

Direct Word <31:0> maps directly into any of the LO’s <DB31:DB0> registers, as directed by the command word. The value of <7:0> must never equal 0x FF, as this signifies a command word, and the LO’s don’t see this as a valid value.

a) For LO1 per the ADI ADF4356 datasheet, program all registers twice from Register 13>>Register 0 with the specified time delays between words.

b) For LO2 and LO3, per the Maxim MAX2871 datasheet, program all registers twice from
Register 5>>Register 0 with the specified time delays between words.

 This command with direct data words allows complete control of any and all of the ADF4356 or MAX2851 registers. However, the price is the increased time to program all registers (twice) with all the delays required. As the Arduino is only passing along the direct data words to the LO registers, in the order that is directed by the PC/Laptop, it is incumbent upon the PC/Laptop to provide the correct order, but the Arduino can easily provide the delays.
The Arduino sketches can program the attenuator and all three LO’s with default values upon power-up and make the use of the General LO commands necessary less frequent. The Specific Commands in section 2 greatly reduces the use of General LO commands to when only necessary.

APPENDIX A. Fractional Synthesis Calculation with MAX2871: Simplified Iterative Algorithm.
As Double-Precision Floating Point Math is used, the PC or Laptop is tasked with this function. Some basic “Pseudo-Code” follows. This algorithm can be implemented in any suitable programming language.

For LO2: Given That LO2[MHz]= (Fref[MHz]/R)*(N+F/M) with Fref[MHz]=60 or 100MHz
and R=2 as our default, and the output divider is set to 1. So Fvco=LO2[MHz]
3000MHz ≤ Fvco ≤ 6000MHz. Fvco is Double or Extended Precision variable.
Integer N Valid Values: 19 ≤ N≤ 4091
Integer F Valid Values: 0 ≤ F ≤ 4095
Integer M Valid Values: 2 ≤ M≤ 4095
Other Variables: Double or Extended Precision: Fract, BestErr, Err

Step 1: Calculate Fpfd=Fref/R
Step 2: Calculate N= INT (Fvco/Fpfd) INT is Integer Function, Converts to Floor()
Step 3: Calculate Fract=(Fvco/Fpfd)-N
Step 4: Initialize BestErr=2^24 or any very high value.

Step 5: Start a For (or a For-Next Loop) to step M from 2 to 4095
Step 6: Calculate F=CINT (Fract*M) CINT is round to even function.
Step 7: Calculate Err=ABS(Fvco-(Fpfd*(N+F/M))) Determines Absolute Value of frequency error.

Step 8: IF Err<BestErr then {
        BestErr=Err
        BestF=F
        BestM=M
        }
Step 9. Back to Step 5 unless M=4095, in which case proceed to Step 10.

Step 10 IF F=M then {
        F= BestF= Best M then {
        BestF=0
        BestM=4095
        N=N+1
        }
Step 11 F=BestF
        M=BestM
        N (N is correct as is).

For LO3: Given That LO3[MHz]= (Fref[MHz]/R)*(N+F/M)/16 with Fref[MHz]=60 or 100MHz
and R=2 as our default, and the output divider is set to 16. So Fvco=LO3[MHz] * 16
3000MHz ≤ Fvco ≤ 6000MHz. Fvco is Double or Extended Precision variable.
The steps follow exactly same as LO2.
Go through Steps 1 through 11 same as for LO2.

APPENDIX B. Sample Program in FreeBasic.

The following program calculates the N,F,M and Output Counter settings for your Reference frequency <60> MHz default for our s, R counter setting <2> default for our purposes, and desired RF output frequency. The user takes these counter settings, then inserts them into the appropriate MAX2871 register, usually as hexadecimal or binary values. This program can be used in a more complete program to program a MAX2871 for single or stepped frequencies. See MAX2871 datasheet for more information.

  'MAX2871_Counter_1.bas for calculating N,F,M and Output Counter Settings
' This program is free software: you can redistribute it and/or modify
' it under the terms of the GNU General Public License as published by
' the Free Software Foundation, either version 3 of the License, or
' (at your option) any later version.
'
' This program is distributed in the hope that it will be useful,
' but WITHOUT ANY WARRANTY; without even the implied warranty of
' MERCHANTABILITY or FITNESS FOR A PARTICULAR . See the
' GNU General Public License for more details.
'
' You should have received a copy of the GNU General Public License
' along with this program. If not, see <https://www.gnu.org/licenses/>.
' Absolutely NO liability is assumed by the author for its use or misuse.

dim shared as integer F,M,R,N,BestM,BestF,Div,DivRange,maxErrHz
dim shared as double Fvco,Fref,Fpfd,Fract,maxErr,Err1,Frf
dim shared as string Ans
declare sub outDiv
Start:
screen 19 :cls : print "FreeBasic Program MAX2871_Counter_1.bas"
print "Counter Settings for Fractional Mode" : print
input "Enter Desired RF output frequency <23.5~6000> MHz ";Frf
outDiv
input "Enter Fref <60> default [MHz] ";Fref : if Fref=0 then Fref=60
input "Enter overall R Counter Setting <2> default ";R : if R=0 then R=2
Fpfd=Fref/R : N=INT ((Fvco)/Fpfd) : Fract=(Fvco/Fpfd)-N

' Loop tries to find best F and M with least frequency error
maxErr=2^24
for M=2 to 4095
F=cint(Fract*M) : Err1=abs( Fvco-(Fpfd*(N+F/M)))

if Err1<maxErr then
  maxErr=Err1
  BestM=M
  BestF=F
  end if
next M
'if BestF=BestM, clear BestF, Make BestM=4095,Increment N
if BestF=BestM then
   BestF=0
   BestM=4095
   N=N+1
   end if

Sample Program in FreeBasic.(Continued)

maxErrHz=cint(maxErr*(1e6)/Div)
print: print "Fvco [MHz]= ";Fvco ;" Output Divider= ";div
print "Fpfd[MHz]= ";Fpfd;" N= ";N;" BestF= ";BestF;" BestM= ";BestM
print "RF Output Frequency Error [Hz]= ";maxErrHz
input "To do another press Enter. To quit enter <q> Enter ";Ans
if Ans="" then goto Start
print "Goodbye,73!" : sleep 2000

sub outDiv 'Enter with Frf leave with Div,Fvco
  for DivRange = 0 to 7 'Determine divider range
    Div=(2^DivRange) : Fvco=Div*Frf
    if (Frf * Div) >= 3000 then goto FoundDiv
  next
FoundDiv:
end sub

APPENDIX C. Table of Default Settings (TBD)

APPENDIX D. Default Arduino Code: Default Register values that can be used prior to issuing General or Specific Commands.
These are Pre-Loaded by the Arduino into the Spectrum Analyzer RF Assembly at power-up or reset.

int ByteA = {0x40}; //16 dB Attenuation Examples: 0x00=0dB 0x40=16dB 0x7F=31.75dB
// Fixed Frequency=MHz, 60 MHz Reference. Codes for 6 total registers
//Digital Lock Detect Enabled, 3.2mA Charge Pump,10 sec Code Refresh

// RF = 146.52 MHz LO1 = 3776.52MHz
unsigned long WordLO1[] = {0x002007E0, 0x0E24DD21, 0x00000002, 0x00000003, 0x32008984, 0x00800025, 0x35012076, 0x060000E7, 0x15596568, 0x0F09FCC9, 0x00C00EBA, 0x0061200B, 0x000015FC, 0x0000000D, 0x02008984}; // R4=0x32008984 Dig Lockdet or R14=0x02008984 TriState

// LO2=3945 MHz Opt Fract Opt; Dig Lockdet MuxOut Fpfd=30 MHz
unsigned long WordLO2[] = {0x00418008, 0x20008011, 0x58009242, 0xF8008003, 0x638FF1DC, 0x80400005, 0x40009242}; //3945MHz PFD=30MHz PowerMax B
// R2=0x58009242 Dig Lockdet or R6=0x40009242 TriState

// LO3=270 MHz Opt Fract Opt; Dig Lockdet MuxOut Fpfd=30 MHz
unsigned long WordLO3[] = {0x00480000, 0x20008011, 0x58009242, 0xF8008003, 0x63CFF104, 0x00400005, 0x40009242};
// R2=0x58009242 Dig Lockdet or R6=0x40009242 TriState